## Slide 1: Innovation from the field GenXchange
Emmanuel ANDRE, MD
Medical Microbiologist
Assistant to the NTP, DRC

## Slide 2: Global context

## Slide 3: South-Kivu Province (DRC) context
6 million hab. Existing network of 100 microscopy centres
Primary indication of Xpert tests
    Context of under detection of TB
    TB REACH project -> Smear negative TB suspects tested with Xpert
    +-20.000 Xpert tests performed to date
Implementation of 10 Xpert machines
  Major cities of the province
  No internet access in 9/10 Xpert sites
  No Laboratory Informatic Systems
  Samples referred from > 50 microscopy laboratories
No prior local experience using molecular biology technologies
Very limited local experience of MDR-TB management
No on-site support from manufacturer for installation of machines, training or maintenance

## Slide 4
Xpert network implemented
3 Xperts in Bukavu (1 million hab)
7 Xperts in cities 20 – 200.000 hab
- 1 only accessible by plane
- 2 only accessible by boat
All requiring 3 – 5 days travel for on-site supervision
Security & logistical issues
1 Xpert moved after year 1 from major mining city (Kamituga) because of impossible adequate supervision

## Slide 6
Significant variation was observed in failure rates both across instruments and over time; furthermore, substantial variation was observed in failure rate in two cartridges lots.
Of the total 108 modules in the study, 34 required replacements on account of various performance related issues; 10 had an overt malfunction and 24 showed degraded performance in test validity rates indicating impending malfunction.

## Slide 7: Development of a surveillance system
Xpert centers (10)
Referring health centers (50 - 100)
Central unit
Early identification of Rif resistant cases
Stock monitoring
Health staff
Patients ( >20.000)
Maintenance

## Slide 8: Short video (2 min)

## Slide 9: Monitoring of activity
Extraction Febr. 2014

## Slide 10: Monitoring of various pre-analytical indicators
Extraction Febr. 2014

## Slide 11: Monitoring of error rates per laboratory
Extraction Febr. 2014

## Slide 12: Monitoring of error rates per module
Extraction Febr. 2014

## Slide 13: Analysis of error rates per site & per module
Extraction Febr. 2014

## Slide 14: Actions performed after observation of abnormal error rates
Centres with general high rate failure
  Additional staff training (conservation of samples, …)
  Calibration (failed) -> replacement
Centres with isolated module dysfunction
  Calibration (failed) -> Replacement
  Time for replacement : +- 2 months
  Actions implemented in March 2014
  Ex : Module 632792 was replaced on 17/3 by new module (625842)
  Direct impact not measured yet (insufficient number of tests performed with new modules)

## Slide 15: Q1 2014 (Jan – March 2014) : 1.706 tests performed
VALID RESULT AVAILABLE (n = 1392 ; 81,6%)
  MTB POSITIVE (n = 139 ; 10%)
    RIF Resistance detected (n = 13 ; 10%)
    RIF INDETERMINATE (n = 19 ; 13,7%)
    No data available (n = 5 ; 3,6%)
    RIF resistance not detected (n = 101 ; 72,7%)
  1253 NEGATIVE (90%)
NO VALID RESULT AVAILABLE (n= 314 ; 18,4%)
  229 ERROR
  79 INVALID
  6 NO RESULT

## Slide 16: Conclusion
Module failure is the major cause of errors in our setting
Cost linked failing modules/cartridges is high
  Direct cost (including transport and importation)
  Additional manpower and cost of on-site travels
Solutions
  Improved quality of modules
  Strict quality controls by manufacturer
  Early detection of abnormal error rates
  Rapid replacement / calibration after identification of problem
Direct issues for DRC
  Handover from « project » to NTP of existing Xpert network showing evident signs of fragility
  Ongoing pressure/support from partners for buying additional Xpert machines without prior consolidation of monitoring system

## Slide 17: Thank you!
IT developper
Benjamin Nyange
NTP/DRC
Dr Georges Bakaswa
Dr Dieudonné Kalumuna
Jean-Paul Chirambiza
Rosette Nyota
Eric Mulume Musafiri
UCLouvain
Prof Michel Delmée
IRD
Ali Habib
Dr Aamir Khan
LSHTM
Dr Olivier le Polain de Waroux
…